Tender No: Neurosurg./11-12/03

Technical Specification of Nerve Monitoring System

- Should have minimum of Eight Channel Monitoring

- Should be capable of monitoring cranial Motor Nerves III,IV,V,VI,VII,IX,X,XI and XII.

- EMG signal must have audio and video representation.

- The equipment should be designed for ease of use with electrode placement screen displayed on the main unit for easy and faster electrode placement.

- Should monitor simultaneously during bipolar cautery.

- Should have Real-time continuous monitoring with ‘Automatic Periodic Stimulation’ TM electrode.

- Should have high sensitivity with reduced interference.

- Should have artifact detection feature to distinguish between artifact and EMG signals.

- Should have provision to connect external keyboard.

- Should have option of surgeon mini screen to display monitoring information on small screen with a provision to mount on an IV pole.

- Should have high input sensitivity with reduced interference.

- Should have high input impedance.

- Should have electrode – checking features.

- Should have frequency response of 100-2000 Hz.

- Preamplifier gain should be at least 107 +/- 4 dB.

- Should have improvised user friendly interface.

- Should have constant current type stimulator.

- Should have capability of deliver current from 0.00 mA to 30 mA

- Should be able to adjust the stimulation range by the lowest stimulation current that can be delivered.

- Should have incrementing probe to adjust stimulation level and print or save data directly by surgeon performing the surgery.

- Stimulating mono polar and bipolar probes of different variety should be available.

- Should be able to save and load custom setting with quick setup.

- Should be able to log EMG activity throughout a procedure for records.

- Should be able to get printable case log report for patient records.

- Should have touch screen for ease of use, with high contrast, digital graphic colour, visible even in darkness.

- Should have color coded channel labeling for easy identification & use.

- Should have USB port for connection with mass storage devices including compact flash drive.

- Should give audio for use with headphone.

- Should have muting option to reduce unwanted signals and artifacts specially cautery noise.

- The monitor should continuously measure electrode impedance and warming should be displayed when out of range.

- Preferable to have option of Patient Simulator of training and education.

- Consumable for 2 and 4 channels each along with stimulating probe should be quoted.

- All other consumable prices should be quoted in option.

- On site guarantee unit and is to be repaired within 48 hours or is to be replaced by alternative equipment.

- Equipment is to be provided at BHU

- Validity not less than 31st March 2012

- Equipment should be of high Quality with excellent repute and Imported

- User list of last 5 years