Tender No: Neurosurg./11-12/04

Technical Specification of Plasma Sterilizer

- System should have cycle temperature less than 55ₒC, should have chamber of 25-30 liter usable volume, and should be mobile. It should run on single – phase 10A/220 volt. Total cycle time less than 30 min, and pre-programmed without any room for human error due to manual programming. The sterilizer detects excess moisture thus eliminating chance of contamination due to residual bioburden. The Sterilant (H202>55%) in a cassette with leak proof indicator, stored at room temperature. Sterilization process involve two equal injection of equal quality of sterlants (H202) and the by-products are non toxic and eco-friendly (water vapor and oxygen), approved by US FDA. System able to document cycle with printout of the cycle parameters and provide one second data for quality control for every second of the cycle. Sterilizer has consumable like cassette: Prefilled H202 with leak proof indicator, chemical indicator strips/tape, biological indicators, instrument trays and matching instrument tray mats, polypropylene and tyvek wrapper for wrapping instrument trays and medical devices, sterilizer quoted model has been present in the market for last 5 years.

- Sterilization Cycles Consumables at least for 500 cycles.

- You are bound to provide onsite / offsite comprehensive guarantee for the items mentioned above for a period of 60 months from the date of satisfactory installation.

- Machine is to be repaired within 48 hours or is to be replaced by alternative machine during repair period.

- Validity not less than 31st March 2012

- Equipment is to be provided at BHU

- User list of last 5 years

- Equipment should be of high Quality with excellent repute and Imported

- Submit the user list of last 5 years