Tender No: Neurosurg./11-12/01

TECHNICAL SPECIFICATION OF “ULTRASOUND SURGICAL ASPIRATOR WITH ENDOSCOPIC ATTACHMENT

Ultrasound Aspirator System should have Fragmentation, Irrigation and Aspiration facilities continuous with low noise and with Endoscopic attachment

- Console should be light weight and portable on trolley with single foot switch.

- Main voltage should be as per Indian standard with perfect class of protection.

- Console and hand piece should be capable of working continuous without any cooling system.

- Should have in build suction system with high vacuum.

- Irrigation rate should be up to 50 ml/min.

- Hand piece should be light weight sterlizable by steam autoclave or ethylene oxide or plasma.

- Hand pieces micro & Macro offered should have choice of outside & inside irrigation; Tip of Hand piece should be reusable: Can be used for long time continuously without cooling and getting heated. Central activation should be through foot switch.

- Macro hand piece, short angled with connecting cable frequency: 25 KHZ, Weight App. 200 g, working length: Approximately 4 cm, Tip DIA: outside Approx. 3 mm, inside approx. 2 mm, Central activation: foot switch.

- Micro suction Pen Hand piece, long angled with connecting cable; Frequency: 35 KHZ, weight: Approx. 100 g, Working length: Approx. 10 cm, Tip DIA: Outside Approx. 2.0-2.5mm, Inside: Approx 1.5-2.0 mm.

- Micro-Pen-Hand Piece, Spatula, Short Spatula Frequency: 35 KHZ, Weight: Approx. 50 g, Working Length: Approx. 30 cm, Spatula Shape.

- Micro Endoscopic Hand piece straight, Frequency: 35 KHZ, Weight: Approx. 100 g, Working Length: Approx. 150 mm, to be used through Karl Storz’s Neuro endoscope with working channel diameters of 3 mm, Activation: Active length approx. 10 – 15 mm, Tip DIA: outside: 2.0 mm, inside approx. 1.5 mm. (To be used through Karl Storz’s Neuroendoscope with working channel diameter of 3mm)

- Consumable should not be expensive and should be freely available.

- Should Magnetorestrictive or Piezoelectric technology based system.

- Resonance frequency of tip should be in range of 20-35 KHZ.

- Should have container for hand pieces for storage & Autoclaving.

- Essential Accessories are be provided free of cost.

- Service is to be provided at site and fault has to be repaired with 48 hours and if not possible, substitute unit is to be provided for smooth functioning of O.T.

- Also quote separately the discount to be offered for buyback of existing surgical aspirator with following details (Equipment may be inspected in Neurosurgery O.T., IMS, BHU).

- Make – SORING (SONOCA)

- Year of Purchase – About 8 years old

- Status of Equipment - not functioning

- Validity not less than March 2012

- Equipment is to be provided at installation site in BHU

- Equipment should be of high Quality with excellent repute and Imported

- User list of 5 year