Technical specifications for “Intravital Real-time Confocal thrombus imaging system"

The confocal system should be of latest and modular technology suitable for fixed and live cell sample imaging. System should be of high sensitivity detection capability to meet various challenging imaging needs of Biological samples including Live Cell imaging applications for FRAP, FRET,FLIP, photo activation and conversion experiments.

The system should be offered with the following configuration:

- Fully Motorized & computer controlled Inverted Fluorescence Research Microscope:

Bright field, Fluorescence and DIC observations with Motorized Z-focus drive with step resolution

of 10 nm or better with dedicated TFT/LCD Touch screen capable of controlling motorized functions

of scope, 6 position motorized FL filter wheel, 6 position motorized DIC nose piece. Motorized XY

Scanning Stage with Universal sample holder for slides, 35 mm Petri dish and Lab tek chambers

with glass bottom cover slips. 12V/100w halogen illumination For BF & DIC and 120W metal Halide

Illuminator with high lifetime of 1500  hours for fluorescence. High Resolution Confocal Grade

objectives 10x/0.30, 20x/0.80, multi immersion  25X/0.8 (water, oil, Glycerol), 40x/1.30 oil,

60/63x/1.40 oil with complete DIC accessories for all objectives.  Band Pass fluorescent filters for

DAPI, FITC/GFP& TRITC/Rhodamine. An imported anti-vibration table for  the complete Microscope

and laser scanning system.

- Complete automated and programmable on stage environmental control chamber (incubator) for live cell time lapse imaging with temperature, CO2 , Humidity control .The system should be able to use 100% CO2 gas  and to provide 5% pre heated CO2 gas to the chamber. All the parameters of the Incubation system should be controlled by software.

- Cooled monochrome CCD Digital camera with 1.4 million pixel chip resolution, 2/3” CCD chip, FireWire IEEE 1394 connectivity, controlled by confocal software for high resolution fluorescence/ DIC imaging.

- Spectral Confocal Laser Scan head with built-in detectors:

High sensitivity confocal laser point scanning and detection unit with built-in spectral detectors for high efficient fluorescence signal collection. Capable of conventional intensity & Spectral based Confocal Imaging for complete visible range. Should be capable of simultaneous imaging of 2 fluorophores and at least 4 in sequential mode. Spectral dispersion of the emission light should be of latest technology with highly efficient spectral separation. Online separation and display of over-lapping emission signals through emission finger printing technique. Motorized & computer controlled continuously variable confocal pinhole with software control. High speed XY Galvo scanner with 360 deg scan rotation with total scan flexibilities of Line, free hand curved line, XY, XYZ, XYZT and XYZT, λ combinations. The laser scanner should have dual scan capability of fast scan for bleaching/photo activation & normal scan for Imaging. Scan resolution at least 2K x 2K for all channels and selected freely down to 4x1 pixels. Scan Zoom range 0.5X to 30X or more and adjustable in steps of 0.1. Scan speed of minimum 5 fps @ 512x512 pixel resolution and should increase to almost 150 fps at 512 X 16. Data acquisition and digitization capability with 8, 12 and 16 bit should be available. An additional transmitted light detector should be offered for bright field and DIC imaging

- Laser Modules:

The system should be offered with the following laser lines to excite the respective fluorochromes. Preference would be given to solid state lasers due to its long life and maintenance free operation.

- 488 nm for Alexa 488, FITC, GFP, Fluo 4, Cy2 fluorochromes.

- 555/559 nm for TRITC, Rhodamine, Texas Red, Cy3, PE, PI fluorochromes.

- 635/639 nm for Alexa 635, DRAQ 5, Cy5 fluorochromes.

- 405 nm for DAPI, Hoechst fluorochromes.

All the lasers should be connected to the scan head through fiber optic cable. All the laser lines should be computer controlled for fast laser switching and attenuation in synchronization with the scanner.

- Control Computer

Latest control computer with Core 2 Duo E8400 processor, RAM 4 GB DDR2-667,HDD: 500GB SATA II,

DVD Super Multi SATA +R/RW, SATI Fire GL V5600 512MB, Giga bit Ethernet, Win 7 OS 64 bit,

USB 2.0, Firewire. Large 30” LCD TFT monitor or dual TFT monitors with 21”.

- System control and imaging software.

Software should be capable of controlling Motorized functions of microscope, scan head control, laser control, scanner control, Image acquisition & processing. Saving of all system parameters with the image for repeatable/reproducible imaging. Capability of Line, curved line, frame, Z-stack, Time series imaging. Photo-activation/conversion, FRAP & FRET imaging capabilities. Ion imaging with online ratio metric imaging and analysis. Standard geometry Measurements like length, areas, angles etc. including intensity measurements. 3D image reconstruction from a Z-stack image series. Co-localization and histogram analysis with individual parameters. Spectral un-mixing and emission fingerprinting technique should be standard feature of the software. Advanced macro for complex time series experiments should be provided.

- Installation and service support.

Bidders should clearly specify the after sales service and application support capabilities. Should provide all pre-installation requirements to have the system installed in ideal room conditions. Provide a detailed list of users of the quoted system in India with contact details.

Optional Accessory to be quoted as optional:

Plan Apochromat 100X Oil/1.4 Objective.